## Table of Contents
Table of Contents
Definitions
Scales
User Needs
Design Inputs
Verification

## Definitions
Term | Definition
Design Control | The end-to-end process of planning, executing, and documenting of the device design process. The design control process includes understanding users' needs all the way through to proving that the final device works as intended. It creates a traceable paper trail connecting every design decision to a user need.
Design Input | A specific, measurable requirement that describes what the device must do or be. If you can't write a pass/fail test for it, it's not a valid design input.
Design Output | The tangible results of the design process, including the device itself, software, drawings, labeling, manufacturing procedures, and documentation such as risk assessments and test reports. Some outputs, like a risk assessment, can also drive new design inputs when they reveal unaddressed hazards or gaps.
Requirement | A singular, verifiable statement defining a specific behavior, characteristic, or constraint that the device must satisfy. Requirements come in different flavors depending on what they describe: functional (what it does), performance (how well it does it), interface (how it connects to users or other systems), safety (what it must never do or always do to prevent harm), usability (how easy and error-resistant it is to use), cybersecurity (how it protects data and resists attack), and regulatory (specific obligations tied to laws or standards).
User Need | A statement of what a user needs to accomplish and why, written from the user's perspective without prescribing a solution. User needs drive everything downstream. Therefore, if they're wrong or incomplete, so is everything that relies on them.
Validation | Confirming that the specification actually reflects what the users need; i.e., did we build the right thing? It involves real or representative users performing real-world tasks, and it's the final check that the whole design control process worked.
Verification | Confirming that you built the device to its specification; i.e., answering the question: did we build it right? Verification is conducted using tests, inspections, analysis, or demonstrations, with no end users involved.

## Scales
User Need Priorities
Priority | Intent | Regulatory Significance
Critical | Required for patient safety, essential clinical function, or regulatory compliance | Failure could cause patient harm or regulatory failure
High | Required for proper clinical operation | Failure impacts clinical workflow
Medium | Important usability or performance feature | Failure causes inconvenience
Low | Enhancement or optional feature | No impact to safety or essential function
User Need Validation Methods
Method | Intent | Example | Regulatory Significance
Clinical Validation | Confirms that the use of the Parkinson's holter monitor ultimately improves treatment outcomes of Parkinson's patient when used by the intended users in clinical settings and patient home environment | Clinician confirms tremor data collected by teh device supports treatment. | Required by FDA 21 CFR 820.30(g) and ISO 13485 §7.3.7 to demonstrate the device is clinically effective and safe
Usability Validation | Confirms that the holter monitor can be used by intended users safely and efficiently, without unacceptable errors or excessive effort. | Clinician can navigates the mobile application and retrieve patient data without confusion. | Required by IEC 62366-1 and FDA Human Factors Guidance to demonstrate safe interaction between user and device
System Validation | Confirms that the complete holter monitoring system fulfills its intended use and user requirements under normal operating conditions. | Full system transmits tremor data from the wearable device via wifi to AWS reliably. | Required by FDA Design Validation requirements and ISO 13485 Design Controls to demonstrate the final production-equivalent system performs correctly
Security Validation | Confirms the whole system protects data confidentiality, integrity, and availability against unauthorized access or tampering | Unauthorized user cannot access patient data. | Required by FDA Cybersecurity Guidance (and §524B of the FD&C Act for premarket submissions) and ISO 13485 to ensure protection of patient data and system integrity
Data Integrity Validation | Confirms that data is accurately stored, transmitted, and retrieved without corruption, loss, or unauthorized modification | Tremor data saved on cloud matches sensor output. | Required by FDA 21 CFR 820 and 21 CFR Part 11 (if applicable), ISO 13485, and IEC 62304 to ensure reliability of medical data
Installation Validation | Confirms that the holter monitor can be properly set up, connected to the network, and function continuously in home or clinical settings for accurate tremor monitoring. | IT administrators can remotely update device firmware and software to address security vulnerabilities and maintain compliance. | Required by ISO 13485 §7.5.3 and FDA Quality System Regulation (21 CFR 820) to ensure correct deployment and operation
Interoperability Validation | Confirms the holter monitoring system accurately and securely exchanges patient data with connected clinical systems (e.g., EHR) under expected integration conditions | Tremor reports generated by the device can be integrated into patient's EHR without significant custom configuration | Required by FDA 21 CFR 820.30 and ONC interoperability provisions; IEC 62304 applies where software interfaces are involved.
Requirement Verification Methods
Method | Intent | Example | Regulatory Significance
Test | Executes the device or system under defined conditions and measures the outcome against a specific, quantitative pass/fail criterion. | Collect and measure latency, packet loss, alert timing, encryption strength, and authentication enforcement. | Use when the requirement is quantitative or behavioral, and the outcome can be directly measured. Most cybersecurity requirements are verified by test.
Inspection | Examines the device, its configuration, documentation, or artifacts visually or analytically to confirm conformance without executing the system. | Review logs, examine configuration files, check display layout, or compare a document against a specification. | Use when conformance can be confirmed by looking rather than running. Inspection does not involve operating the device.
Demonstration | Operates the device to show that a function works as intended, typically without precise measurement, relying on observable correct behavior. | Review report generation, printing, and retrieval of historical records. | Use when the requirement describes a workflow or capability that is best confirmed by observing it in action rather than measuring a specific value. Common for usability and functional requirements.
Analysis | Evaluates conformance through calculation, modeling, or systematic review of data rather than direct device operation or inspection of artifacts. | Use threat modeling, vulnerability assessment, software architecture review, code review, or compliance evaluation. | Use when direct testing is impractical or insufficient on its own. Analysis produces a documented rationale rather than a measured result.

## User Needs
ID | Description | User Type | Rationale/Use Case | Priority | Asset Map | Validation Method
UN1 | Patients need the wearable device to accurately and reliably detect and quantify tremor episodes in real time to support continuous Parkinson's symptom monitoring. | Patient | Continuous tremor monitoring during daily activity | Critical | A1 | Clinical Validation
UN2 | Patients need the wearable device to be lightweight and unobtrusive so it can be worn comfortably during extended daily use. | Patient | Long-term wearability and patient compliance | Critical | A1 | Usability Validation
UN3 | Patients need the device to operate continuously to support uninterrupted monitoring. | Patient | Uninterrupted daily monitoring | High | A1 | System Validation
UN4 | Patients need to view their data to be aware of their own disease progression. | Patient | Patient self-awareness to disease progression | High | A3 | Usability Validation
UN5 | Patients need their personal health data to be protected from unauthorized access to preserve privacy and comply with applicable regulations. | Patient | Patient data privacy and regulatory compliance | Critical | A1, A4, A6, A7 | Security Validation
UN6 | Clinicians need access to real-time and historical tremor data to assess disease progression and conduct retrospective analysis. | Clinician | Disease progression monitoring and treatment optimization | Critical | A1, A3, A6, A7 | Clinical Validation
UN7 | Clinicians need tremor severity scores and movement metrics to be displayed clearly and accurately to support clinical interpretation. | Clinician | Clinical interpretation of tremor data | Critical | A3, A5 | Usability Validation
UN8 | Clinicians need to associate tremor data with the correct patient to support accurate patient identification and record management. | Clinician | Patient identification and record integrity | Critical | A1, A3, A4, A6, A7 | Usability Validation
UN9 | Clinicians need to generate patient tremor reports to support clinical documentation and care coordination. | Clinician | Clinical documentation and reporting | High | A3 | Usability Validation
UN10 | Clinicians need accelerometer data to be reliably transmitted from the wearable device to the cloud without data loss to support remote patient monitoring. | Clinician | Remote monitoring of ambulatory Parkinson's patients | Critical | A1, A4 | Data Integrity Validation
UN11 | Clinicians need the system to reliably classify tremor events using validated signal processing algorithms to minimize false positives and negatives. | Clinician | Accurate tremor classification for clinical decision-making | Critical | A5 | Clinical Validation
UN12 | Clinicians need to operate the tremor monitoring system without specialized technical training to support routine patient monitoring workflows. | Clinician | Routine patient monitoring workflows | Critical | A1, A3 | Usability Validation
UN13 | Caregivers need to receive timely notifications when a patient experiences a significant tremor episode so they can provide assistance. | Caregiver | Caregiver awareness and timely patient assistance | High | A3, A8 | System Validation
UN14 | Caregivers need to view a summary of the patient's recent tremor activity to monitor daily symptom burden. | Caregiver | Daily symptom burden monitoring by non-clinical users | High | A3, A5 | Usability Validation
UN15 | Caregivers need the mobile or web application to be simple and easy to navigate without medical or technical expertise. | Caregiver | Accessible interface for non-clinical users | High | A3 | Usability Validation
UN16 | IT administrators need the system to maintain comprehensive audit logs of all access and data modification events to support compliance monitoring. | IT Administrator | Regulatory compliance and audit traceability | Critical | A7 | Security Validation
UN17 | IT administrators need the ability to remotely update device firmware and software to address security vulnerabilities and maintain compliance. | IT Administrator | Secure remote software update management | High | A1, A3, A5 | Installation Validation
UN18 | IT administrators need the system to integrate with existing clinical or hospital infrastructure without requiring significant custom configuration. | IT Administrator | Seamless clinical environment integration | Medium | A1, A6, A7 | Interoperability Validation

## Design Inputs
ID | Description | Allocation | Category | UN Ref(s) | Verification Method
DI1 | The system shall display real-time symptom metrics to clinicians with a refresh rate of no less than 1 frame per second. | Software | Functional | UN1, UN4, UN6, UN14 | Test
DI2 | The system shall acquire raw accelerometer signals with a sampling rate of at least 20 Hz | Hardware | Functional | UN1, UN6 | Test
DI3 | The system shall acquire raw accelerometer signals with an resolution of at least 2.87 cm/second2 per increment. | Hardware | Functional | UN1, UN6 | Test
DI4 | The longest dimension of the wearable component (excluding fastening) shall not exceed 20 cm | Hardware | Usability | UN2 | Test
DI5 | The weight of the wearable component shall be less than 60 g | Hardware | Usability | UN2 | Test
DI6 | The battery of the wearable component shall last more than 24 hours under normal operating conditions without recharging. | Hardware | Usability | UN3 | Test
DI7 | The system shall associate all acquired accelerometer data with a unique patient identifier prior to storage or transmission. | Software | Functional | UN8 | Inspection
DI8 | The system shall allow retrieval and display historical symptom recordings for a specified patient. | Software | Functional | UN6, UN7, UN14 | Demonstration
DI9 | The system shall enable clinicians to generate and export symptom reports in PDF format. | Software | Functional | UN9, UN11 | Demonstration
DI10 | The system shall support printing of symptom reports via a connected printer. | Software | Functional | UN9, UN11 | Demonstration
DI11 | End-to-end latency from signal acquisition to symptom metric display shall not exceed 500 ms under normal operating conditions. | System | Performance | UN1 | Test
DI12 | The system shall transmit patient data between assets with a packet loss rate of no greater than 3% under normal home wireless conditions. | System | Performance | UN10 | Test
DI13 | The system shall generate an audible and visual alert for the caregiver within 10 seconds of detecting a clinically significant "off" period or device fault. | System | Safety | UN11, UN13 | Test
DI14 | The system shall maintain continuous symptom monitoring and alerting functions during transient wireless disconnection of up to 30 seconds. | System | Safety | UN3 | Test
DI15 | The system shall provide a status indicator showing connectivity state, battery level, and active alarms visible from the main clinician view. | Software | Interface | UN3 | Inspection
DI16 | The system shall expose a FHIR R4-compliant API for exchange of symptom observations with connected EHR systems. | Software | Interface | UN18 | Test
DI17 | The system shall allow a trained clinician or caregiver to register a new patient and begin symptom monitoring in no more than 3 user interactions. | Software | Usability | UN12, UN15 | Inspection
DI18 | The system shall maintain an audit log of all access to patient records, including user ID, timestamp, and action performed, retained for a minimum of 3 years. | Software | Regulatory | UN16 | Inspection
DI19 | The system shall enforce role-based access control (RBAC) such that clinicians, IT administrators, and service accounts are granted only the minimum permissions required for their function. | Software | Cybersecurity | UN5 | Test
DI20 | The system shall enforce multi-factor authentication (MFA) for all IT administrator accounts at both the login endpoint of the application’s administrative web dashboard and for SSH remote maintenance access to the embedded devices prior to granting access to any system configuration or maintenance functions. | Software | Cybersecurity | UN5 | Test
DI21 | All patient data transmitted between the Raspberry Pi, mobile app, and AWS backend systems shall be encrypted using TLS 1.2 or higher with cipher suites that provide authenticated encryption and forward secrecy (e.g., AES-GCM or ChaCha20-Poly1305 with ECDHE). | Software | Cybersecurity | UN5 | Test
DI22 | All patient data stored on the system shall be encrypted at rest using AES-256 or an equivalent algorithm in a secure mode of operation (e.g., AES-GCM or AES-XTS). Insecure modes of operation such as ECB shall not be used. | Software | Cybersecurity | UN5 | Test
DI23 | The holter monitor shall restrict all network communications except for outbound MQTT over TLS on port 8883 for outbound data transmission and SSH on port 22 for remote maintenance. The cloud system shall restrict all network communications except for MQTT over TLS on port 8883 to receive telemetry data and HTTPS on port 443 for data transfer with authorized mobile applications. | System | Cybersecurity | UN1, UN5, UN17 | Test
DI24 | The holter monitor shall transmit MQTT telemetry data only to the authorized AWS IoT Core endpoints. | System | Cybersecurity | UN1, UN5 | Test
DI25 | The holter moniter shall allow SSH remote maintenance connections only through an authenticated TailScale VPN tunnel. | System | Cybersecurity | UN5, UN17 | Test
DI26 | The system shall provide a mechanism to roll back a failed software or firmware update to the last known good version. | Software | Cybersecurity | UN1, UN5, UN17 | Test
DI27 | The system shall log all authentication events, access control decisions, and security-relevant configuration changes to a tamper-evident audit log. | Software | Cybersecurity | UN16 | Test
DI28 | The mobile and desktop app shall implement session timeout, locking out inactive authenticated sessions after no more than 15 minutes with no user input. | Software | Cybersecurity | UN5 | Test
DI29 | The system shall reject all login attempts for 15 minuts after 3 failed login attempts to prevent brute forcing the credentials | Software | Cybersecurity | UN5 | Test
DI30 | The system shall require all passwords to be a combination of letters, numbers, and symbols that is 10-15 characters long. | Software | Cybersecurity | UN5 | Test

## Verification
ID | DI Refs | Verification Activity | Method | Pass Criteria | Test Date | Tester | Status | Artifacts Location
V1 | DI18 | Inspect audit log contents and retention policy configuration. | Inspection | Log captures user ID, timestamp, and action; retention ≥3 years confirmed | 2026-04-24 | Rex Wang | Pass |  | Note: Access to patient data stored in S3 bucket baymax-tremor-archive-8011 are logged in a separate bucket baymax-access-log which retains entries for 3 years.
V2 | DI8 | Using the mobile app, retrieve and display historical symptom recordings for a pre-loaded test patient with known data across multiple sessions. | Demonstration | Historical recordings are retrieved and displayed for the specified patient; displayed values match the known stored records with no missing sessions. |  |  | Pass
V3 | DI9 | Using the mobile app, generate and export a symptom report in PDF format for a specified patient and date range. | Demonstration | A valid, readable PDF file is generated and exported containing the correct patient name, date range, and symptom data; PDF opens without error in a standard viewer. |  |  | Pass
V5 | DI15 | Inspect the main clinician view of the mobile app and confirm that connectivity state, battery level, and active alarm indicators are present and visible without scrolling or navigating to another screen. | Inspection | All three indicators (connectivity state, battery level, active alarms) are visible on the main clinician view without any additional navigation; each indicator updates to reflect a simulated state change within the display refresh period. |  |  | Pending | TBD
V6 | DI23 | Perform a network port scan and service enumeration on the holter monitor and cloud system to identify all open ports and running services; confirm that only the ports and services specified in the allowlist are accessible. | Test | Only documented ports allowed in DI23 are open; no unexpected services running | 2026-04-22 | Rex Wang | Pass |  | Note: nmap shows only opened port on raspberry pi is ssh
V7 | DI24 | Attempt to connect holter to a test MQTT broker. | Test | The holter monitor rejects or cannot establish connection to non-authorized endpoints in all test cases | 2026-04-24 | Rex Wang | Pass |  | Note: unit test showing device refuse to connect to unauthenticated brokers
V8 | DI25 | Attempt SSH connection from a device not connected to TailScale. | Test | The connection fails and no SSH access granted in all test cases | 2026-04-21 | Rex Wang | Pass
V9 | DI26 | Simulate failed update; verify rollback to last known good version without data loss. | Test | System restores to prior version after simulated update failure; patient data and configuration intact | 2026-04-24 | Rex Wang | Pass |  | Note: current implementation checks SHA256 hash and python syntax. New known good version can be added by run.sh -promote. Future implementation can check git integrity.
V10 | DI27 | Review audit log contents after performing authentication, access control, and configuration change events. | Test | All specified event types present in log with correct fields; log file integrity mechanism verified | 2026-04-21 | Rex Wang | Pass
VER-17 | DI17 | Ask a trained clinician or caregiver to register a new patient and begin symptom monitoring from the mobile app home screen; an observer counts every discrete user interaction (tap or confirmation). | Inspection | Task is completed in ≤3 user interactions from the home screen to active monitoring for the new patient. |  |  | Pending | TBD
VER-21 | DI21 | Capture network traffic from the mobile app to the backend during a representative monitoring session using a packet analyser; inspect TLS version and negotiated cipher suite. | Test | All captured patient-data traffic from the mobile app uses TLS 1.2 or higher; negotiated cipher suites provide authenticated encryption and forward secrecy (e.g., ECDHE-AES-GCM or ECDHE-ChaCha20-Poly1305); no plaintext patient data is observed in any captured packet. |  |  | Pending | TBD
VER-22 | DI22 | Inspect local storage on the mobile device (e.g., app sandbox, SQLite database, cached files) and review the mobile app's encryption implementation in source code or configuration. | Test | All patient data stored locally by the mobile app is encrypted with AES-256 or equivalent in a secure mode (GCM or XTS); ECB mode is absent; encryption is confirmed via both source/config review and direct storage inspection. |  |  | Pending | TBD
VER-28 | DI28 | Authenticate on the mobile app, leave the session idle for 15 minutes, then attempt to perform an authenticated action without re-authenticating. | Test | The mobile app automatically locks the session after ≤15 minutes of inactivity; any subsequent authenticated action is blocked until the user re-authenticates successfully. |  |  | Pending | TBD
VER-29 | DI29 | Submit 3 consecutive failed login attempts with incorrect credentials on the mobile app login screen; immediately attempt a 4th login; wait for the lockout period to expire and attempt again with correct credentials. | Test | Login is blocked after 3 failed attempts; the 4th attempt is rejected within the 15-minute lockout window; a successful login with correct credentials is possible only after the full 15-minute lockout period has elapsed. |  |  | Pending | TBD
VER-30 | DI30 | On the mobile app, attempt to set passwords that violate the policy (< 10 chars, > 15 chars, letters only, numbers only, no symbols); then submit a compliant password (10–15 chars with letters, numbers, and at least one symbol). | Test | All non-compliant passwords are rejected with an informative error message; compliant passwords (10–15 characters containing letters, numbers, and at least one symbol) are accepted. |  |  | Pending | TBD